| Block | Specification | Comments |
| --- | --- | --- |
| BioZ Excitation Frequency | DC to 500 kHz | Programmable DDS |
| Excitation Current | 0.032-256 µA RMS (default), optional up to 1.28 mA | Smartwatch optimized |
| Receiver Architecture | Analog I/Q Demodulation | Similar concept to MAX30009 |
| Input Filter | EMI + Programmable HPF | Remove DC electrode offset |
| INA Gain | Programmable 1×–10× | High CMRR (>103 dB) |
| I/Q Mixer | Analog Passive Mixer | Quadrature demodulation |
| LPF Cutoff | Programmable 0.6–2 kHz | Depends on application |
| PGA Gain | 1×–10× | After LPF |
| ADC Type | 20-bit Sigma-Delta | Separate I and Q ADC |
| ADC Resolution | 20 bits | Raw conversion |
| ADC OSR | 8-2048 | Programmable |
| ADC ODR | 16–4096 SPS | Programmable |
| 20-bit Digital LPF | 000: Bypass, 001: 0.005 × Sample Rate, 010: 0.02 × Sample Rate, 011: 0.08 × Sample Rate, 100: 0.25 × Sample Rate | Noise reduction |
| 20-bit Digital HPF | 00: Bypass, 01: 0.00025 × Sample Rate, 10: 0.002 × Sample Rate | Remove DC drift |
| Magnitude Calculation | √(I²+Q²) | Hardware/DSP |
| Phase Calculation | atan2(Q,I) | Hardware/DSP |
| Output Resolution | 20-bit I, 20-bit Q | Before magnitude calculation |
| Interface | SPI | Register programmable |

# Low-Power Bioimpedance Measurement Using Analog IQ Demodulation

## Tissue Impedance Model

Biological tissue exhibits both resistive and reactive electrical characteristics. Consequently, the tissue impedance is represented as a complex quantity rather than a purely resistive value. The impedance can be expressed in either rectangular or polar form as:

Z=R+jX

or equivalently

Z=∣Z∣ejθ
where R is the tissue resistance, which represents the real component of the impedance. Xis the tissue reactance, which represents the imaginary component arising primarily from cell membrane capacitance. ∣Z∣ is the impedance magnitude and θ is the impedance phase angle.

The objective of the bioimpedance measurement system is to accurately determine both the magnitude and phase of the tissue impedance, since these quantities provide valuable physiological information about the tissue under test.

## Excitation Current

To measure the tissue impedance, a known sinusoidal current is injected into the body using a Direct Digital Synthesis (DDS) current source. The excitation current is given by

iexc(t)=I0sin⁡(ωt)
where I0 is the excitation current amplitude, ω=2πfexc is the excitation frequency and fexc=50kHz is the excitation frequency.

## Tissue Voltage Response

According to AC Ohm's Law:

V=IZ

the injected current flowing through the complex tissue impedance produces a voltage response

vtissue(t)=I0∣Z∣sin⁡(ωt+θ)

The amplitude of the measured voltage is therefore

Av=I0∣Z∣
while its phase shift is

ϕ=θ

This relationship demonstrates that the tissue converts the known excitation current into a sinusoidal voltage whose amplitude is proportional to the impedance magnitude and whose phase shift directly corresponds to the impedance phase.

## Analog IQ Demodulation

Directly measuring the amplitude and phase of a high-frequency sinusoidal signal is both power-intensive and computationally expensive. To enable a low-power implementation, the proposed architecture employs analogue IQ demodulation.

The measured tissue voltage

v(t)=I0∣Z∣sin⁡(ωt+θ)

is simultaneously multiplied by two locally generated reference signals that are exactly 90° apart. These reference signals are sin⁡(ωt) and sin⁡(ωt).

The resulting outputs are commonly referred to as the in-phase (I) and quadrature (Q) channels. After multiplication, a low-pass filter removes the high-frequency mixing products, leaving only the DC components that contain the desired impedance information.

In-Phase (I) Channel

The received voltage is multiplied by the in-phase reference sin⁡(ωt) gives

v(t)sin⁡(ωt)=I0∣Z∣sinωt+θsin(ωt)
Using trigonometric identity

sinAsinB=12[cos⁡(A-B)-cos⁡(A+B)]
produces

=I0∣Z∣2[cos⁡θ-cos⁡(2ωt+θ)]
The second term oscillates at twice the excitation frequency and is removed by the low-pass filter. Consequently, the I-channel output becomes:

I=I0∣Z∣2cos⁡θ
The I-channel therefore represents the real (in-phase) component of the complex impedance.

Quadrature (Q) Channel

Similarly, the received signal is multiplied by the quadrature reference cos⁡(ωt) gives

v(t)sin⁡(ωt)=I0∣Z∣sin⁡(ωt+θ)cos(ωt)

Using trigonometric identity

sinAcosB=12[sin⁡(A+B)+sin⁡(A-B)]
results in:

=I0∣Z∣2[sin⁡(2ωt+θ)+sin⁡θ]
After the low-pass filter removes the high-frequency term, the remaining output becomes:

Q=I0∣Z∣2sin⁡θ

The Q-channel therefore represents the quadrature component of the measured tissue impedance.

## Complex Measurement

The outputs of the two demodulation channels are combined to form a complex measurement:

M=I+jQ

Substituting the expressions for the I and Q channels yields

M=I0∣Z∣2(cos⁡θ+jsin⁡θ)
Using Euler’s identity,

ejθ=cos⁡θ+jsin⁡θ
the measurement simplifies to

M=I02∣Z∣ejθ
Since:

∣Z∣ejθ=Z
the final expression becomes

M=I02Z

This important result shows that the output of the analog IQ demodulator is directly proportional to the complex tissue impedance.

## Ideal Impedance Recovery

If the excitation current is perfectly known and the analog front-end introduces no gain or phase errors, the tissue impedance can be recovered directly from the measured I and Q components as

Z=2I0M
or

Z=2I0(I+jQ)

This represents the ideal operating condition of the measurement system.

## Error Model

In an integrated circuit implementation, several circuit non-idealities introduce gain and phase errors into the measured signal. These include excitation current inaccuracies, mixer conversion gain, PGA gain variation, ADC gain error, component mismatch, and temperature drift.

Rather than modelling each error independently, all non-idealities can be combined into a single complex calibration constant (K). The measured signal is therefore expressed as

M=KZ
where (K) represents the overall gain and phase response of the complete measurement chain.

## Calibration Using a Precision Resistor

To determine the calibration constant, a precision resistor with a known resistance (Rcal) is connected to the measurement path.

Since the calibration resistor is purely resistive,

Zcal=Rcal
The corresponding complex measurement is

Mcal=Ical+jQcal
Substituting into the measurement model gives

M=KZ
Therefore, the calibration constant is

K=McalRcal

Once this constant has been determined, all subsequent measurements can be corrected for the gain and phase errors of the analog front-end.

## Unknown Tissue Measurement

After calibration, the resistor is disconnected and the tissue is reconnected to the measurement system.

The measured complex output is

Mx=Ix+jQx
Using the practical measurement model,

Mx=KZx
Substituting the previously calculated calibration constant,

K=McalRcal
gives

Zx=MxK

Finally,
Zx=RcalMxMcal

or equivalently,

Zx=RcalIx+jQxIcal+jQcal

This equation forms the basis of the calibration algorithm. Rather than relying on the absolute gain of the analog circuitry, the unknown impedance is obtained by comparing the measured tissue response with the response produced by a known reference resistor.

## Extraction of Resistance and Reactance

Although the complex equation above completely describes the unknown impedance, many biomedical applications require the resistance and reactance to be reported separately.

To separate the real and imaginary components, the denominator is multiplied by its complex conjugate.

Multiplying both the numerator and denominator by the conjugate gives:

Zx=RcalIx+jQx)(Ical-jQcalIcal2+Qcal2

Expanding the numerator produces separate expressions for the real and imaginary components.

The tissue resistance is therefore

Rx=RcalIxIcal+QxQcalIcal2+Qcal2

Similarly, the tissue reactance becomes

Xx=RcalQxIcal-IxQcalIcal2+Qcal2
These equations provide the real and imaginary components of the measured tissue impedance while automatically compensating for the analog front-end gain and phase errors determined during calibration.

## Magnitude and Phase Calculation

Once the resistance and reactance have been calculated, the impedance magnitude and phase angle can be obtained using standard complex-number relationships.

The impedance magnitude is calculated as

∣Zx∣=Rx2+Xx2
The impedance phase angle is calculated as

θx=tan-1XxRx
The magnitude indicates the overall opposition of the tissue to the injected current, while

the phase angle reflects the relative contributions of the resistive and capacitive properties of the biological tissue.

These two parameters represent the final outputs of the bioimpedance measurement system and can be used in a wide range of biomedical applications, including body composition analysis, respiration monitoring, bioimpedance spectroscopy, impedance cardiography, and other physiological monitoring techniques.

## Summary

The complete analogue IQ demodulation process can be summarised as follows:

- A precise sinusoidal current is injected into the biological tissue.

- The tissue converts the excitation current into a voltage whose amplitude and phase are determined by the complex tissue impedance.

- The received voltage is multiplied by two orthogonal reference signals to obtain the in-phase (I) and quadrature (Q) components.

- Low-pass filters remove the unwanted high-frequency mixing products, leaving only the DC values containing the impedance information.

- The I and Q outputs are combined to form a complex measurement.

- A precision calibration resistor is measured to determine the overall gain and phase response of the analogue front-end.

- The calibration constant is used to compensate for circuit non-idealities, allowing the resistance, reactance, impedance magnitude, and impedance phase of the unknown tissue to be accurately calculated.

This calibration approach provides high measurement accuracy while maintaining a low-power architecture, making it well suited for wearable and battery-powered bioimpedance measurement systems.

## Example of Analogue IQ Demodulation

To illustrate the analog IQ demodulation process, consider a simple example where a

100 Ω precision resistor is connected to the measurement system instead of biological tissue. Since the load is purely resistive, the impedance has no imaginary component, and the expected phase shift is 0°. This example demonstrates how the I and Q channels recover the impedance using synchronous detection.

A sinusoidal excitation current with an amplitude of 1.28 mA and a frequency of 50 kHz is generated by the DDS current source. The excitation current is

iexct=I0sinωt=0.00128sin2π×50000×t

As the excitation current flows through the 100 Ω resistor, a voltage is developed across the load according to AC Ohm's Law, V=IZ. The resulting voltage is:

vloadt=0.128sin2π×50000×t

Since the load is purely resistive, the voltage is perfectly in phase with the excitation current and no phase shift is introduced.

The received voltage is first applied to the in-phase mixer, where it is multiplied by a reference signal that is perfectly synchronized with the excitation current. The mixer output is

vIt=0.128sin2π×50000×tsin2π×50000×t

Using the trigonometric identity:

sinAsinB=12cosA-B-cosA+B

gives

vIt=0.1282cos0-cos4π×50000×t=0.064-0.064cos4π×50000×t

The first term is the desired DC component, while the second term is an AC component at 100 kHz, which is twice the excitation frequency. The output is then passed through a low-pass filter that removes the 100-kHz component, leaving only the DC value I=0.064 V. Thus, the I channel successfully extracts the in-phase component of the measured impedance.

The same received voltage is simultaneously applied to a second mixer. In this channel, the reference signal is shifted by 90°, producing the quadrature measurement. The mixer output is

vQt=0.128sin2π×50000×tcos2π×50000×t

Using the trigonometric identity:

sinAcosB=12sinA+B+sinA-B

gives

vQt=0.1282sin4π×50000×t+sin0=0.064sin4π×50000×t

Unlike the I-channel, this signal contains only a 100-kHz component and no DC term. After passing through the low-pass filter, the high-frequency component is removed completely, resulting in Q=0 V. This is the expected result because a purely resistive load produces no quadrature component.

The outputs of the I and Q channels are combined to form the complex measurement M=I+jQ. From the IQ demodulation analysis,

M=I02Z

Rearranging the equation gives

Z=2I0M

Substituting the excitation current and the measured complex output,

Z=20.001280.064+j0=100Ω

The recovered impedance is identical to the actual resistor value, confirming that the analog IQ demodulation process accurately reconstructs the impedance of a purely resistive load. As expected, the in-phase channel contains the entire signal, while the quadrature channel is zero because there is no phase difference between the excitation current and the measured voltage.

In practical bioimpedance measurements, biological tissue is not purely resistive. The cell membrane behaves as a capacitor, causing the measured impedance to become complex and introducing a phase difference between the injected current and the measured voltage.

To demonstrate this behavior, consider a 100 Ω resistor connected in parallel with a 1 nF capacitor. This simple RC network represents a basic approximation of biological tissue.

Since the resistor and capacitor are connected in parallel, it is convenient to calculate the admittance of the network first. The admittance is

Y= 1100+j2π×50000×10-9=0.01+j0.00031416 S

The impedance is the reciprocal of the admittance. Therefore,

Z= 99.901-j3.138 Ω

The impedance magnitude is

Z=99.9012+3.1382≈99.95

The impedance phase angle is

tan-1-3.13899.901=-0.0314 radians≈-1.8°

The negative phase angle indicates that the circuit exhibits capacitive behavior. Consequently, the measured voltage lags the injected current by approximately 1.8°.

The excitation current remains unchanged from the previous example. Similarly, the voltage developed across the load becomes

vloadt=0.128×99.95sin2π×50000×t-0.0314=0.127936sin2π×50000×t-0.0314

The received voltage is multiplied by the in-phase reference signal.

vIt=0.127936sin2π×50000×t-0.0314sin2π×50000×t

After multiplication and low-pass filtering, only the DC component remains.

I=0.1279362cos-1.8=0.063968 V

Since the phase shift is very small, almost all of the signal energy remains in the I-channel.

The same voltage is simultaneously multiplied by the quadrature reference signal.

vQt=0.127936sin2π×50000×t-0.0314cos2π×50000×t

After the low-pass filter,

Q=0.1279362sin-1.8=-0.002008 V

Unlike the purely resistive example, the Q channel is no longer zero. The small negative value represents the capacitive component introduced by the parallel capacitor.

The complex measurement is

M=0.063968-j0.002008

And therefore,

Z=20.001280.063968-j0.002008=99.9-j3.138 Ω

The recovered impedance closely matches the theoretical impedance of the RC network, demonstrating that the analog IQ demodulator accurately measures both the resistive and reactive components of the load.

The analog IQ demodulator always reports the measured impedance as a series equivalent impedance

Zs=Rs+jXs

However, many biological tissue models represent the tissue as a resistor in parallel with a capacitor. Therefore, the measured series impedance must be converted into an equivalent parallel resistance and capacitance.

The equivalent parallel admittance is

Y=1Z

Substituting the measured impedance,

Yp=1Rp+jωCp

So, we need to convert series Z to parallel Y:

Y=1Rx+jXx=Rx-jXxRx+jXxRx-jXx=RxRx2+Xx2-jXxRx2+Xx2

1Rp=RxRx2+Xx2

Rp=Rx2+Xx2Rx

ωCp=-XxRx2+Xx2

In this case

Rp=99.9012+3.138299.901=9990.05684599.901=99.999568 Ω≈100Ω

Cp=3.1389990.056845×2π×50k=9.9985059878×10-10≈1nF

The calculated parallel resistance and capacitance closely match the original circuit values, confirming that the conversion is correct.

Next, consider a resistor and capacitor connected in series. The series impedance is

Zs=100-j3183.1Ω

The impedance magnitude is

Z=1002+3138.12≈3184.67

The phase angle is

tan-1-3138.1100=-1.539 radians≈-88.2°

Since the capacitive reactance is much larger than the resistance, the phase angle approaches −90°, indicating that the circuit behaves almost like a pure capacitor.

The voltage across the load is

vloadt=0.00128×3184.67sin2π×50000×t-1.539=4.076sin2π×50000×t-1.539

After IQ demodulation,

I=4.0762cos-1.539=0.064 V

Q=4.0762sin-1.539=-2.037 V

Recovering the impedance gives

M=0.064-j2.037 V

Z=20.001280.064-j2.037 =100-j3182.8 Ω

The recovered impedance differs from the theoretical value by only approximately 0.3 Ω in the imaginary component. This small discrepancy is entirely due to rounding during the intermediate calculations and is not caused by the IQ demodulation process itself. Finally, the equivalent series capacitance is

Cs=13183.1×2π×50k=9.999996424×10-10≈1nF

This confirms that the analog IQ demodulation architecture accurately reconstructs both series and parallel RC networks while preserving the resistive and reactive components of the impedance.

# Gain Limitation Due to Offset and Noise

Although increasing the analogue gain improves the impedance resolution, the gain cannot be increased indefinitely. Amplifier offsets and electronic noise consume part of the ADC input range, reducing the available headroom before clipping occurs.

To prevent saturation,

Vsig+Vostotal+Vnoisepeak≤VREF

where

Vsig=ZLIexcAtotal

The total input-referred offset referred to the ADC input is

Vostotal=VosINAAINAAMIXERALPFAPGA+VosMIXERALPFAPGA+VosLPFAPGA+VosPGA

Rearranging the previous equation gives the maximum allowable analog gain

Atotalmax=VREF-Vostotal-VnoisepeakIexc_rms×ZL

# Measurable Impedance Range

For a given excitation current and analogue gain, the maximum measurable impedance is limited by the ADC full-scale input.

The maximum measurable impedance is

Zmax=VREFAtotalIexc

The smallest practical impedance is limited by the ADC quantisation step. To provide sufficient design margin, a minimum signal amplitude of approximately 10 LSB is assumed.

Therefore,

Zmin=10VLSBAtotalIexc

These equations define the impedance range over which the desired resolution can be achieved.

# Impedance Resolution

Unlike repeatability, which is limited by the analogue noise floor, impedance resolution is limited by the finite quantisation step of the ADC. It represents the smallest discrete change in impedance that can be distinguished by the measurement system.

Since the excitation current is constant, a small impedance variation produces a proportional change in the measured voltage. Therefore, the minimum detectable impedance change is directly related to one Least Significant Bit (LSB) of the ADC.

The ADC LSB is

VLSB=VREF2N

where N is the effective number of bits (ENOB), here. Referring this voltage back to the electrode input,

Vres_in=VLSBAtotal

since

Vres_in=Iexc_rms×∆Zres

the impedance resolution becomes

∆Zres=VREF2N×Iexc_rms×Atotal

where

Atotal=AINAAMIXERAPGAALPF

For the proposed architecture, ALPF=1.

This equation shows that the impedance resolution improves with

- higher ADC resolution,

- higher excitation current, and

- higher analogue gain

Conversely, reducing any of these parameters increases the minimum detectable impedance step.

Substituting the maximum allowable gain into the impedance resolution equation gives

∆Zres=VREF2N×Iexc_rms×VREF-Vostotal-VnoisepeakIexc_rms×ZL=VREF×ZL2N×VREF-Vostotal-Vnoisepeak

For convenience,

Verror=Vostotal+Vnoisepeak

therefore,

∆Zres=VREF×ZL2N×VREF-Verror

This equation directly relates the achievable impedance resolution to the ADC resolution, measurement range and available voltage headroom.

Rearranging the previous equation,

VREF-Verror=VREF×ZL2N×∆Zres

Verror=VREF-VREF×ZL2N×∆Zres

Verror=VREF1-ZL2N×∆Zres

This equation indicates the maximum offset and noise that can be tolerated while still achieving the desired impedance resolution. Since the available ADC input headroom must remain positive (VREF > Verror), the ADC must satisfy,

ZL2N×∆Zres<1

or equivalently,

ZL∆Zres<2N

This inequality establishes the fundamental requirement that the ADC must provide sufficient quantization levels to resolve the desired impedance change over the specified impedance range.

# Required ADC ENOB

The ADC Least Significant Bit (LSB) is

VLSB=VREF2ENOB

The corresponding voltage change produced by the target impedance resolution is:

VLSB=∆ZresIexcAtotal

Equating these two expressions gives:

VREF2ENOB=∆ZresIexcAtotal

Substituting

Atotalmax=VREF-VerrorIexc_rms×ZL

results in

2ENOB=VREF∆ZresIexcAtotal=VREF∆ZresIexcVREF-VerrorIexc_rms×ZL=ZL∆ZresVREFVREF-Verror

Taking the base-2 logarithm,

ENOB=log2ZL∆Zres+log2VREFVREF-Verror

This equation provides valuable insight into the factors governing the ADC resolution. The first term log2ZL∆Zres depends solely on the measurement range and the required impedance resolution. It represents the theoretical minimum ADC resolution assuming an ideal analog front-end. The second term, log2VREFVREF-Verror, accounts for the reduction in available ADC input range caused by amplifier offsets and electronic noise. As the available voltage headroom decreases, additional ADC resolution is required to maintain the same impedance resolution.

Therefore, achieving a finer impedance resolution can be accomplished by:

- increasing the ADC ENOB,

- reducing the measurement range,

- reducing the analogue front-end offset and noise, or

- increasing the available ADC input headroom.

Table 1 summarizes the required ADC ENOB and the corresponding maximum allowable analog gain for different combinations of impedance range, target impedance resolution, excitation current, and available ADC headroom. Several important observations can be made from the results.

First, the required ENOB is primarily determined by the ratio of the maximum impedance to the target impedance resolution. As expected, measuring larger impedance ranges or achieving finer impedance resolution requires a higher ADC resolution. For example, increasing the maximum impedance from 100 kΩ to 1 MΩ increases the required ENOB from approximately 18 bits to 21 bits when maintaining a resolution of 0.5 Ω. The required ENOB is dominated by the ratio between the maximum impedance and the target impedance resolution. Variations in ADC reference voltage and available voltage headroom have a comparatively smaller effect, typically changing the required ENOB by only a fraction of a bit over the operating conditions considered.

Second, improving the target impedance resolution from 0.5Ω to 0.25Ω increases the required ENOB by approximately 1 bit. This agrees with the theoretical relationship, since halving the quantization step requires twice as many ADC codes.

Third, increasing the available ADC headroom by reducing the offset and noise (smaller Verror) allows a larger analog gain before the ADC saturates. For example, when Verror = 0.2 V, the allowable gain is significantly higher than when Verror = 0.5 V. Consequently, reducing the analog front-end offset directly relaxes the gain design constraints.

Fourth, the excitation current is selected according to the expected impedance range. Higher body impedances require lower excitation currents to keep the electrode voltage within the allowable input range, whereas lower body impedances can be driven with larger currents to improve the signal-to-noise ratio.

Finally, the proposed operating range of 25 Ω to 5 kΩ requires only 14–16 bits ENOB to achieve an impedance resolution of 0.5 Ω. Allowing additional margin for process, voltage, and temperature (PVT) variations, an ADC with 15–16 bits ENOB is sufficient for the proposed bioimpedance measurement system.

Table 1 also summarizes the maximum allowable analog gain for different body impedances and excitation currents.

The results indicate that the required analog gain decreases as the excitation current increases. This is because a larger excitation current generates a proportionally larger differential voltage across the load, requiring less amplification to fully utilize the ADC input range.

For the intended measurement range, the calculated maximum gain is generally below 100. Therefore, an overall analog gain of approximately 100 is sufficient to cover all operating conditions. Increasing the gain beyond this value provides little improvement in impedance resolution while significantly tightening the offset and noise requirements of the analog front-end.

For example, when measuring a 25 Ω load with an excitation current of 256 μA RMS, the differential input voltage is

Vsig = 25 × 256 μA = 6.4 mV

An analogue gain of approximately 78 amplifies this signal to approximately 499.9 mV, nearly utilizing the available ADC input range while maintaining sufficient margin to prevent clipping.

Table 1 ADC ENOB calculation for maximum allowable analog gain and for different combinations of impedance range, and target impedance resolution

| ENOB | ZL(Ω) | ∆Zres(Ω) | VREF | Verror | Iexc | Atotalmax |
| --- | --- | --- | --- | --- | --- | --- |
| 21 | 1M | 0.5 | 1V | 0.2V | 32nA | 25 |
| 17.7 | 100k | 0.5 | 1V | 0.2V | 320nA | 25 |
| 20.98 | 1M | 0.5 | 2V | 0.2V | 32nA | 56.25 |
| 17.66 | 100k | 0.5 | 2V | 0.2V | 320nA | 56.25 |
| 22 | 1M | 0.25 | 1V | 0.2V | 32nA | 25 |
| 18.71 | 100k | 0.25 | 1V | 0.2V | 320nA | 25 |
| 21.2 | 1M | 0.5 | 1V | 0.5V | 32nA | 15.625 |
| 17.9 | 100k | 0.5 | 1V | 0.5V | 320nA | 15.625 |
| 20 | 500k | 0.5 | 1V | 0.2V | 32nA | 50 |
| 16.71 | 50k | 0.5 | 1V | 0.2V | 320nA | 50 |
| 16.92 | 50k | 0.5 | 1V | 0.5V | 320nA | 31.25 |
| 15.4 | 20k | 0.5 | 1V | 0.2V | 640nA | 62.5 |
| 14.4 | 10k | 0.5 | 1V | 0.2V | 1.28µA | 62.5 |
| 15.4 | 10k | 0.25 | 1V | 0.2V | 1.28µA | 62.5 |
| 13.4 | 5k | 0.5 | 1V | 0.2V | 3.2µA | 50 |
| 13.6 | 5k | 0.5 | 1V | 0.5V | 3.2µA | 31.25 |
| 14.4 | 5k | 0.25 | 1V | 0.2V | 6.4µA | 25 |
| 12.1 | 2k | 0.5 | 1V | 0.2V | 12.8µA | 31.25 |
| 12.3 | 2k | 0.5 | 1V | 0.5V | 12.8µA | 19.53 |
| 13.2 | 2k | 0.25 | 1V | 0.5V | 12.8µA | 19.53 |
| 6.95 | 25 | 0.25 | 1V | 0.5V | 256µA | 78.125 |

# Offset Voltage Budget

The allowable headroom loss, Verror, consists of the total analog offset and a design margin reserved for noise and other uncertainties.

Verror=Vostotal+Vnoisepeak

As derived from the Impedance Repeatability section, the total input-referred system noise must be less than 1.28 μV rms to achieve a repeatability of 1 mΩ. Assuming the maximum analog gain of 100, this input-referred noise becomes 128 μVrms when referred to the ADC input. This value is approximately 0.13 mV, which is several orders of magnitude smaller than the available voltage headroom. Consequently, the noise contribution is negligible compared with the offset voltage and can be ignored when establishing the offset budget. The total offset referred to the ADC input is therefore

Vostotal=VosINAAINAAMIXERALPFAPGA+VosMIXERALPFAPGA+VosLPFAPGA+VosPGA

Substituting the selected gains AINA = 10, AMIXER = 0.636, ALPF = 1, APGA = 10 gives

Vostotal=VosINA×10×0.636×1×10+VosMIXER×1×10+VosLPF×10+VosPGA

Vostotal=VosINA×63.6+VosMIXER×10+VosLPF×10+VosPGA

## Offset Budget Allocation

For the proposed design, an allowable headroom loss of 0.2 V is assumed. To account for residual noise, process and temperature variations, modelling inaccuracies, and additional implementation uncertainties, a 10 mV design margin is reserved. Consequently, for an allowable headroom loss of 0.2 V, the remaining 190 mV is allocated to the offset budget of the analog front-end. The offset is allocated among the analog blocks according to their expected contribution to the total system offset.

- INA: 50%

- Mixer: 20%

- LPF: 20%

- PGA: 10%

Thus,

VosINA×63.6=0.5×190mV

VosMIXER×10=0.2 ×190mV

VosLPF×10=0.2 ×190mV

VosPGA=0.1 ×190mV

Then

VosINA=1.49 mV

VosLPF=VosMIXER=3.85 mV

VosPGA=19 mV

These values establish the maximum allowable offset for each analog building block while ensuring that the total offset referred to the ADC input remains within the available voltage headroom.

# Noise Allocation

The total input-referred noise of the bioimpedance measurement system is determined by the combined noise contribution of every analog building block. Since the signal chain is divided by the IQ mixers, the dominant noise mechanisms differ before and after the demodulation stage.

The Analog High-Pass Filter (AHPF) and the Instrumentation Amplifier (INA) operate at the excitation carrier frequency (fexc). These blocks generate broadband thermal noise centered around the carrier frequency. During synchronous demodulation, the IQ mixers translate both the desired signal and the surrounding broadband noise to baseband. Consequently, the broadband noise generated before the mixer is subjected to double-sideband (DSB) noise folding, increasing its contribution to the baseband noise.

In contrast, the Low-Pass Filter (LPF) and the Programmable Gain Amplifier (PGA) operate entirely at baseband (typically DC to a 4Hz). Since these circuits process very low-frequency signals, their performance is dominated by flicker (1/f) noise rather than broadband thermal noise.

The analog front-end consists of two distinct noise regions separated by the IQ mixers. The analog high-pass filter (AHPF) and instrumentation amplifier (INA) operate at the excitation frequency (fexc) and therefore contribute broadband noise around the carrier. During synchronous demodulation, this high-frequency noise is translated to baseband by the mixer through double-sideband noise folding. In contrast, the low-pass filter (LPF) and programmable gain amplifier (PGA) operate directly at baseband, where low-frequency flicker (1/f) noise becomes the dominant contributor.

To evaluate the overall system performance, the noise contribution of every circuit block is referred to the electrode input. Assuming that all noise sources are statistically independent, the total input-referred noise is obtained using the root-sum-square (RSS) method.

The total input-referred noise is therefore

Vnsys2=VnAHPF2+VnINA2+VnMIXERBB21AINA2+VnLPFBB21ALNAAMIXER2+VnPGABB21AINAAMIXERALPF2

Substituting the selected gain values

Vnsys2=VnAHPF2+VnINA2+VnMIXERBB21102+VnLPFBB2110×0.6362+VnPGABB2110×0.636×12=VnAHPF2+VnINA2+VnMIXERBB21100+VnLPFBB2140.45+VnPGABB2140.45

The scaling factors indicate that the noise generated by the mixer, LPF and PGA is attenuated when referred to the electrode input because these stages are preceded by analog gain. Conversely, the AHPF and INA are located at the front of the signal chain, so their noise appears directly at the system input without attenuation. As a result, the first-stage circuits generally dominate the overall input-referred noise budget.

An additional effect occurs during the demodulation process. The broadband noise produced by the AHPF and INA is centered around the excitation frequency. When this noise passes through the IQ mixers, noise components located above and below the carrier are translated simultaneously to baseband. This phenomenon is known as double-sideband (DSB) noise folding and approximately doubles the contribution of broadband white noise after demodulation.

If the mixers are implemented using square-wave switching, the switching waveform contains odd harmonics of the excitation frequency. These harmonics also translate high-frequency noise into the baseband, producing additional harmonic noise folding. Consequently, the effective baseband noise depends not only on the white noise generated by the front-end circuits but also on the harmonic content of the mixer switching waveform.

Therefore, minimising the broadband noise of the AHPF and INA is particularly important, as any noise generated before the mixers is translated to baseband and cannot be removed by the subsequent low-pass filtering stage. In contrast, the LPF and PGA should be designed with low flicker noise to preserve the measurement accuracy at very low frequencies, where physiological impedance variations are observed.

## Passive Chopping Mixer Signal and Noise Analysis

### Signal Conversion Gain

Before analyzing the noise performance of the passive chopping mixer, it is first necessary to determine its signal conversion gain. Unlike a conventional analog multiplier, a passive chopping mixer multiplies the input signal by a square-wave switching function whose amplitude alternates between +1 and −1. The Fourier series of the switching waveform is

mt=4πcosωct-13cos3ωct+15cos5ωct+…

The body impedance consists of a large static component and a small physiological variation

Zt=ZL+∆Zcosωbt

where ωb is the physiological modulation frequency.

With a sinusoidal excitation current

Iexc(t)=Iexccosωct

the electrode voltage becomes

Vsigt=IexccosωctZL+∆Zcosωbt

using the trigonometry identity

cosAcosB=12[cosA-B+cos⁡(A+B)]

gives

Vsigt=IexcZLcosωct+Iexc∆Z2cosωct-ωbt+Iexc∆Z2cosωct+ωbt

The resulting spectrum consists of three components:

- Carrier at ωc with amplitude IexcZL

- Lower Sideband (LSB) atωc-ωb with amplitude Iexc∆Z2

- Upper Sideband (USB) at ωc+ωb with amplitude Iexc∆Z2

Since the subsequent low-pass filter removes all higher-order switching harmonics, only the fundamental component of the square wave contributes to the desired signal demodulation. The LSB is therefore multiplied by

4πcosωct

Thus
VLSBt=Iexc∆Z2cosωct-ωbt×4πcosωct

After multiplication and low-pass filtering,

VLSB_BBt=Iexc∆Z24π×12cos-ωbt=Iexc∆Zπcosωbt

Repeating the same analysis for the upper sideband gives

VUSB_BBt=Iexc∆Z24π×12cosωbt=Iexc∆Zπcosωbt

Unlike noise, the upper and lower signal sidebands originate from the same physiological impedance variation. Consequently, they are perfectly phase coherent and add directly in voltage.

Vtotal_BBt=VLSB_BBt+ VUSB_BBt=2Iexc∆Zπcosωbt

The original physiological envelop before ethe mixer is:

Venvt=Iexc∆Zcosωbt

The total baseband signal recovered after the mixer is:

Voutt=2πIexc∆Zcosωbt

Therefore, the signal voltage conversion gain of a chopping mixer is 2π≈0.636. This conversion gain is used throughout the subsequent system-level analysis.

### Double Sideband Folding

The broadband thermal noise generated by the AHPF and INA behaves differently from the physiological signal. Thermal noise is random and occupies a continuous spectrum around every harmonic of the excitation frequency. When this broadband noise is multiplied by the square-wave switching function, every odd harmonic translates noise from both its upper and lower sidebands to the same baseband frequency. This process is known as double-sideband (DSB) noise folding. Unlike the physiological signal, the upper and lower noise sidebands are statistically independent because they originate from random thermal motion. Consequently, their voltages cannot be added directly. Instead, their powers must be summed.

Consider a single odd harmonic of the square-wave switching function. The amplitude of the kth harmonic is

Ak=4kπ

Assume that the input white-noise power spectral density around this harmonic is

en2 (V2/Hz)

To determine how much of this noise is translated to baseband, consider a 1-Hz bandwidth around one sideband. Within this bandwidth, the RMS noise voltage satisfies

Vn2=en2×1

Since the RMS value of a sinusoid equals its peak amplitude divided by √2, an equivalent time-domain representation of this narrowband noise is

nsidebandt=2Vncosωsbt+φ

where φ is a random phase. The kth harmonic of the switching waveform is:

mkt=Akcoskωct

The mixer output is therefore

yt=2Vncosωsbt+φAkcoskωct

Using trigonometric identity

cosAcosB=12[cosA-B+cos⁡(A+B)]

gives

yt=22VnAkcosωsbt-kωct+φ+cosωsbt+kωct+φ

The low-pass filter removes the high-frequency sum term, leaving only the difference-frequency component

yBBt=22VnAkcosωBBt+φ

where

ωBB=ωsb-kωc

The peak amplitude of the baseband sinusoid is therefore

22VnAk

The average power of a sinusoid having peak amplitude C is

P=C22

Substituting the peak amplitude gives

PBB,k=0.5×22VnAk2=Vn2Ak24

Since the calculation was performed in a 1-Hz bandwidth,

Vn2=en2

Therefore,

PBB,k=en2Ak24

This result represents the baseband noise power contributed by one sideband of the kth switching harmonic. However, every harmonic has both an Upper Sideband (USB) and a Lower Sideband (LSB). Because these two noise sidebands are statistically independent, their powers add after demodulation.

Therefore,

PBB,k=2×en2Ak24=en2Ak22

substituting the amplitude (Ak=4kπ)

PBB,k=en224kπ2=en28k2π2

The total baseband white-noise power is obtained by summing the contribution of every odd switching harmonic,

Pn,out=en28π21+132+152+…

The series converges to a constant value of π28. Thus,

Pn,out=en2

This important result shows that although each individual harmonic contributes only a fraction of the original noise power, the combined contribution of all odd harmonics exactly reproduces the original white-noise power spectral density at the mixer output.

Unlike the noise, the desired physiological signal experiences a voltage conversion gain of only 2π. Therefore, when the mixer output noise is referred to the electrode input, the equivalent input-referred noise power becomes

en, inref2=Pn,out2π2=en2π24=2.4674en2

Thus, the broadband thermal noise generated before the passive chopping mixer appears approximately 2.47 times larger when referred to the electrode input. Therefore, when the thermal noise of the AHPF and INA passes through a chopping mixer and is evaluated at the electrodes, its power is multiplied by 2.4674. This increase is not caused by additional noise generation within the mixer itself, but rather by the finite signal conversion gain of the passive mixer combined with double-sideband noise folding.

### Noise Budget Allocation

From the repeatability analysis, the maximum allowable input-referred system noise is

Vnsys=1.28 μVrms

Or

Vnsys2=1.28 μVrms2=1.6384 pVrms2

To ensure that no individual circuit dominates the overall noise performance, the available noise power is allocated across the analog front-end according to the expected contribution of each block according to Table 2.

Table 2 Noise power allocation

| Block | Noise Power Allocation |
| --- | --- |
| INA | 60% |
| AHPF | 10% |
| Mixer | 10% |
| LPF | 10% |
| PGA | 10% |

The corresponding input-referred noise power budgets are

INA: 0.98304 pVrms2

AHPF: 0.16384 pVrms2

Mixer: 0.16384 pVrms2

LPF: 0.16384 pVrms2

PGA: 0.16384 pVrms2

These values provide the maximum allowable input-referred noise contribution of each circuit block and form the basis for the transistor-level noise design of the analog front-end.

### Pre-mixer Allocation (AHPF and INA)

The AHPF and INA are located before the passive chopping mixer. Consequently, their broadband thermal noise is translated to baseband through double-sideband (DSB) noise folding. As derived in the previous section, this increases their input-referred noise power by a factor of 2.4674.

The allocated noise budgets in Table X represent the maximum allowable integrated input-referred noise power. To determine the corresponding spot noise density specification for each circuit block, the allocated noise power must be related to its white-noise spectral density.

Assuming the white-noise density is approximately constant over the measurement bandwidth, the integrated noise power is

Vrms2=enINA2×ENBW

where ENBW is the equivalent noise bandwidth of the baseband low-pass filter. Including the DSB noise-folding factor gives

Allocated Noise Power = enINA2×2.4674 ×ENBW

The equivalent noise bandwidth is determined by the selected baseband low-pass filter configuration, which limits the integrated white noise to approximately 4 Hz. Therefore

Allocated Noise Power = enINA2×2.4674 ×4= enINA2×9.8696

The allocated input-referred noise power for the INA is 0.98304 pVrms2 therefore

0.98304 pVrms2=enINA2×9.8696

enINA≈315 nV/Hz

Similarly, the allocated noise power for the AHPF is 0.16384 pVrms2 therefore

0.1638 pVrms2=enAHPF2×9.8696

enAHPF≈128 nV/Hz

If INA and AHPF are merged,

0.1638 pVrms2+0.98304 pVrms2=enINAAHPF2×9.8696

enINAAHPF≈340 nV/Hz

### Post-mixer Baseband Allocation

Unlike the AHPF and INA, the mixer, LPF, and PGA operate entirely at baseband (DC to 4 Hz). Since these blocks are located after the passive chopping mixer, their noise is not affected by double-sideband (DSB) noise folding. Instead, their noise is simply referred to the electrode input by dividing it by the gain of all preceding stages.

The mixer follows the INA, and its noise power is:

VnMIXERBB21100=0.16384 pVrms2

VnMIXERBB≈4.04 μVrms

The LPF follows both the INA and the passive mixer. Therefore:

VnLPFBB2140.45=0.16384 pVrms2

VnLPFBB≈2.57 μVrms

The PGA is the final stage and follows the INA, mixer, and LPF. Therefore,

VnPGABB2140.45=0.16384 pVrms2

VnPGABB≈2.57 μVrms

# System Dynamic Range

Since the excitation current is maintained at a constant amplitude, the measured voltage is directly proportional to the load impedance according to AC Ohm's Law. Consequently, the required measurement dynamic range can be calculated directly from the ratio of the maximum measurable impedance to the smallest detectable impedance variation. The dynamic range is defined as

DR=20log10ZbaseZnoise

Where Zbase is the massive static baseline body impedance. Znoise is the RMS impedance repeatability, which represents the smallest impedance variation that can be reliably resolved above the measurement noise floor. For the MAX30009, when the body impedance is 25 Ω, the specified repeatability is 0.8 mΩ rms. Therefore

DR=20log10250.0008=89.9 dB

Similarly, for a body impedance of 100 kΩ, the specified repeatability is 2.53 Ω rms.

DR=20log101000002.53=91.9 dB

These calculations show that the measurement system requires approximately 90–92 dB of dynamic range across the supported impedance range.

If this dynamic range were achieved solely by an ideal ADC, the corresponding Effective Number of Bits (ENOB) can be estimated using

DR=6.02×ENOB+1.76

Substituting DR = 92 dB gives ENOB=15 bits. However, the MAX30009 specifies an effective resolution of approximately 17 bits.

# Repeatability

Repeatability is a measure of the system's ability to produce consistent impedance measurements when the load remains unchanged. In a noisy measurement system, repeated measurements of the same static impedance will fluctuate because the measured voltage is corrupted by electronic noise. As a result, poor repeatability limits the smallest impedance variation that can be reliably detected.

The minimum resolvable impedance, ∆Zrms, represents the system's noise floor expressed in the impedance domain. Since the bioimpedance measurement system injects a constant sinusoidal current, the corresponding input-referred voltage noise can be calculated directly using AC Ohm's Law.

The total input-referred noise voltage must satisfy

Vnsys≤Iexc_rms×∆Zrms

where

- Vnsys is the total input-referred RMS noise voltage.

- Iexc_rms is the RMS excitation current.

- ∆Zrms is the specified impedance repeatability.

Assuming

Iexc_rms=1.28mArms

∆Zrms=1 mΩ

the maximum allowable input-referred noise is

Vnsys≤1.28m×1m=1.28 μVrms

Therefore, the total input-referred noise of the analog front-end, including the instrumentation amplifier, mixers, filters, PGA, and ADC, must be less than 1.28 μV rms to achieve an impedance repeatability of 1 mΩ rms.

# Maximum and Minimum Measurable Impedance

The measurable impedance range is determined by the available excitation voltage, excitation current, and the maximum input range of the receiver. Since

Vexc=Iexc×ZL

the maximum measurable impedance is limited by the maximum allowable excitation voltage, while the minimum measurable impedance is limited by the maximum available excitation current. Table 3 summarizes the impedance range achievable using current-mode excitation.

Table 3 Impedance range

| AC Differential Voltage | Iexc_rms | ZL |
| --- | --- | --- |
| 1V | 16nA | 62.5 MΩ |
| 1V | 3.2 µA | 312.5 kΩ |
| 1V | 320 µA | 31.25 kΩ |
| 1V | 1.28mA | 781 Ω |
| 64mV | 1.28mA | 50 Ω |
| 8mV | 320 µA | 25 Ω |

For the target impedance range of 25 Ω to 100 kΩ, the required excitation voltage varies from only 8 mV to 320 mV. After the maximum programmable analogue gain of 100, these voltages become 0.8 V and 32 V, respectively. Consequently, the highest analogue gain cannot be used for the entire impedance range, and the INA/PGA gain must be reduced as the measured impedance increases to avoid saturating the ADC.

Using the available excitation-current settings, the system is capable of measuring impedances from 25 Ω up to 62.5 MΩ, which is significantly wider than the intended operating range.

In voltage-mode excitation, the DAC must source the current required by the load. Therefore,

IDAC=VexcZL

For the smallest supported load of 25 Ω, the required DAC current is shown in Table 4.

Table 4 Required DAC current for different load impedance

| Amax | Excitation Voltage (mVp) | ZL | Required DAC Current |
| --- | --- | --- | --- |
| 20 | 50 | 25Ω | 2mA |
| 20 | 50 | 39.0625Ω | 1.28mA |
| 10 | 100 | 78.125 Ω | 1.28mA |
| 4 | 250 | 195.3125 Ω | 1.28mA |
| 2 | 500 | 31.25MΩ | 16 nA |

Supporting 25 Ω with a 50mV excitation requires a 2 mA DAC current, which exceeds the available maximum stimulation current of 1.28 mA. Consequently, the smallest impedance that can be driven at this voltage using the available current source is approximately

Zmin=50m1.28m=39.1Ω

At the opposite extreme, the minimum stimulation current of 16 nA can support impedances up to 31.25 MΩ with a 500mV excitation. Although this is approximately one-half of the theoretical 62.5 MΩ achievable with a 1 V excitation, it is still more than two orders of magnitude larger than the target measurement range. Therefore, the available DAC current range is sufficient for voltage-mode excitation over the intended operating range, although very low impedance measurements require either a reduced excitation voltage or a higher maximum output current. It is worth noting that the maximum excitation current of 1.28 mA is required to satisfy patient safety requirements for excitation frequencies above 16 kHz.

# The Mathematical Model of IQ Mismatch

The incoming high frequency BioZ voltage signal is

Vin=Acosωt+φ

In an ideal zero-IF IQ demodulator, the local oscillator (LO) generates two perfectly orthogonal reference signals with identical gain. However, in real analog circuits, there will be a gain mismatch and phaser error between I and Q channels:

LOI=2cosωt

LOQ=-2(1+ε)sinωt+θe

I Channel:

VI=Acosωt+φ×2cosωt=Acosφ+cos2ωt+φ

Q Channel:

VQ=Acosωt+φ×-2(1+ε)sinωt+θe=A(1+ε)sinφ-θe-sin2ωt+φ+θe

The LPF completely removes the high frequency terms:

VI=Acosφ

VQ=A(1+ε)sinφ-θe

The resistive cross-talk problem

In BioZ, the true resistance is

R= Acosφ

and the true reactance is

X= Asinφ

Let’s expand the VQ equation:

VQ=A1+εsinφ-θe=A1+εsinφcosθe-sinθecosφ

B substituting real R and X into this equation:

VQ=1+εXcosθe-Rsinθe

Because of the phase error, a portion of the resistance R leaks into the quadrature measurement via - Rsinθe term.

Measured Magnitude Error

The measured magnitude is calculated as

Ameas=VI2+VQ2

Substituting the LPF outputs:

Ameas=Acos2φ+(1+ε)2sin2φ-θe

Based on the R and X definitions, the fully expanded baseband voltages are:

Vsig(t)=IexcZsinωct+θ

Resistance: R=Zcosθ

Reactance: X=Zsinθ

In an ideal zero-IF IQ demodulator, the LO generates two perfectly orthogonal signals. However, in reality the LO suffers from gain mismatch (∆A) and phase error mismatch (∆φ):

LOI(t)=2(1+∆A)sinωct+∆φ

LOQ(t)=2(1-∆A)cosωct-∆φ

The mixer multiplies the input signal by the I-channel LO:

VmixI=IexcZsinωct+θ2(1+∆A)sinωct+∆φ

Using

2sinA sinB=cos⁡(A-B)-cos⁡(A+B)

A=ωct+θ

B=ωct+∆φ
VmixI=IexcZ(1+∆A)cosθ-∆φ-cos2ωct+∆φ+θ

The LPF completely removes the high frequency 2ωct term and we are left with the DC baseband component:

VI=IexcZ(1+∆A)cosθ-∆φ

Using

cosA-B=cosAcosB+sinA sinB

VI=Iexc1+∆AZcosθcos∆φ+Zsinθsin∆φ=Iexc1+∆ARcos∆φ+Xsin∆φ

The mixer multiplies the input signal by the Q-channel LO:

VmixQ=IexcZsinωct+θ2(1-∆A)sinωct-∆φ

Using

2sinA cosB=sinA-B+sin⁡(A+B)

A=ωct+θ

B=ωct-∆φ
VmixQ=IexcZ(1-∆A)sinθ+∆φ+sin2ωct-∆φ+θ

The LPF completely removes the high frequency 2ωct term and we are left with the DC baseband component:

VQ=IexcZ(1-∆A)sinθ+∆φ

Using

sinA+B=sinAcosB+cosA sinB

VQ=Iexc1-∆AZsinθcos∆φ+Zcosθsin∆φ=Iexc1-∆AXcos∆φ+Rsin∆φ

We apply small angle approximation:

cos∆φ≈1

sin∆φ≈∆φ

∆A×∆φ≈0

VI≈IexcR1+∆A+X∆φ

VQ≈IexcX1-∆A+R∆φ

Zmeas≈R21+2∆A+X21-2∆A+4RX∆φ

Using a Tylor series expansion

u+ε≈u+ε2u

And assuming R>>x thus:

Zmeas≈Z+R∆A+2X∆φ

## Impact on Impedance Resolution

Term IexcR∆φ in VQ is a static DC offset injected directly into the Q channel. If base resistance (R) is larger than X, even a phase error of 1° = 0.017 radians causes a massive voltage drop.

APGA_Q×IexcX+R∆φ≤VREF

APGA_Q=VREFIexcX+R∆φ

Because the denominator is artificially inflated by the leaked resistance R∆φ the maximum allowable gain is forced down. As proven previously, lower gain widens LSB step size, reducing the resolving power (∆Zres) of the Q-channel.

## Impact on SNR and ENOB

The error vector magnitude is defined by the extra terms in our magnitude expansion:

Zerror=R∆A+2X∆φ

SNR=20log10ZtrueZerror+VnoisesysIexc

ENOB=20log10ZtrueZerror+VnoisesysIexc-1.766.02

## Impact on impedance repeatability

Static IQ mismatch creates a constant offset (which can be calibrated out), drifting IQ mismatch destroys repeatability.

Numerical example:

Excitation current: 1mA

True Resistance: 100 Ω

True Reactance: 2 Ω

True Magnitude: 100.02 Ω

True Phase: 1.145°

Gain Mismatch: 1%

Phase Error: 1°

VI=1mA 1+0.01100cos1+2sin1=1.01×99.98+0.0349=101.015 mV

VQ=1mA1-0.012cos1+100sin1=0.99×1.999+1.745=3.706 mV

Zmeas=101.0152+3.70621mA=101.08 Ω

Error=101.8Ω-100.02Ω=1.06Ω which is significantly larger than the target accuracy. MAX30009 achieves 0.2% gain matching for I vs Q channel.

## IQ Mismatch Mitigation:

### Calibration Phase:

During a calibration phase, the IC measures Rcal (which has zero true reactance, X=0). Any voltage that appears in the Q channel during this test is strictly the resistive crosstalk (R∆φ). Any deviation from the expected I-channel voltage is the gain error (∆A).

### High Frequency LO generation

To reduce the phase error, at the digital side, clocks running at 2 times of the desired excitation frequency is recommended. This clock is fed into a D-Flip-Flop (DFF).

### Layout:

Common-centroid is highly recommended to guarantee that any thermal gradients across the chip or oxide thickness gradients from the silicon wafer manufacturing affect both the I and Q channels equally, canceling out the mismatch.

# Oscillator Frequency Stability Requirement

The excitation oscillator determines the carrier frequency used for bioimpedance measurement. Its frequency stability affects measurement in two independent ways. First, because biological tissue exhibits reactive behavior, any change in excitation frequency changes the actual impedance of the tissue. Second, oscillator drift alters the phase response of the analog front-end, producing IQ mismatch after calibration. Both effects must therefore be considered when specifying the oscillator frequency accuracy.

## Frequency Drift and Tissue Reactance

Human tissue and cell membranes behave as biological capacitors. Capacitive reactance Xc is strictly dependent on frequency

Xc=-1ωexcC=-12πfexcC

Since the reactance depends directly on the excitation frequency, any oscillator drift changes the measured impedance even when the patient's physiological state is perfectly constant.

Differentiating the reactance with respect to frequency gives

dXcdfexc=ddfexc-12πfexcC=12πCfexc2=-Xcfexc

For a small frequency variation,

∆Xc=-Xc∆fexcfexc

This simple result shows that the reactance error is proportional to both the true tissue reactance and the fractional oscillator drift. For example, if the body reactance is −5 Ω (Xc=-5Ω) and the oscillator drifts by 1% (∆fexcfexc=0.01),

∆Xc=5×0.01=50mΩ

Therefore, to limit the reactance error below 100 mΩ, the oscillator frequency drift should remain below approximately 2%. This corresponds to a maximum frequency variation of 4 kHz at a 200 kHz excitation frequency and 1 kHz at a 50 kHz excitation frequency.

## Frequency Drift on IQ Mismatch:

Oscillator drift also changes the phase response of the analog high-pass filter. Since the calibration coefficients are obtained at the nominal excitation frequency, any subsequent frequency drift introduces a residual phase error that appears as IQ mismatch. The cutoff frequency of AHPF is defined as:

fc=12πRC

The complex transfer function is defined:

Hjω=jωRC1+jωRC

To find the phase shift, we must subtract the phase angle of the denominator from the phase angle of the numerator.  The numerator is purely imaginary (jωRC). The angle of a purely imaginary number is exactly 90°. The denominator is a complex number and its angle is:

Denominator angle =tan-1ffc

So, the phase shift at any given excitation frequency is:

θ=π2-tan-1fexcfc

Using the standard trigonometric identity: π2-tan-1x=tan-11x, we arrive at the simplified equation:

θfilterfexc=tan-1fcfexc

During calibration, the excitation frequency is fexc0, producing a fixed phase shift

θ0=tan-1fcfexc0

This θ0 is measured and is used to for calibration. At this exact moment, the system is perfect. There is zero phase error, and therefore zero resistive crosstalk into the Q channel. Because of the heat the oscillator frequency drifts to a new frequency:

fexc_new=fexc0+ ∆fexc

Because the frequency changed, the phase shift of the AHPF also changed. To find out exactly how much the phase shifted, we must take the derivative of the phase equation with respect to the excitation frequency.

ddxtan-1u=11+u2dudx

u=fcfexc0

The derivative of u with respect to fexc is:

ddfexcfcfexc-1=-fcfexc-2

Thus:

dθdfexc=11+fcfexc2-fcfexc-2=fexc2fexc2+fc2-fcfexc-2=-fcfexc2+fc2

Using a linear approximation, the actual drift in phase is simply the derivative (the slope) multiplied by the physical frequency drift:

∆θ≈dθdfexc∆fexc=-fcfexc2+fc2∆fexc

The analog hardware just physically shifted the signal phase by ∆θ. However, the digital calibration matrix is completely blind to this drift. It is still strictly subtracting the original θ0.

If we plug this ∆θ back into the fundamental IQ mismatch equation. As proven earlier, any uncompensated phase error between the signal and the mixer LO forces a portion of the massive resistance to bleed into reactive Q-channel:

VQerror≈IexcR∆θ=IexcR(-fcfexc2+fc2∆fexc)

Xerror=IexcR∆θIexc=R∆θ

This equation shows that the reactive measurement error increases linearly with body resistance. Consequently, oscillator frequency stability becomes increasingly important as the measured impedance increases. We need to have ultra-stable resistive calibration load. The firmware designer should rapidly switch the excitation and the measurement path away from the patient and into calibration R for a few milliseconds.

## Numerical examples

Excitation frequency: 100KHz

AHPF cutoff: 10KHz

Oscillator drift: 1% (1KHz)

Actual Resistance: 100Ω

Excitation Current: 1mA

∆θ≈dθdfexc∆fexc=-1041052+1042103=-0.00099 radians=-0.056°

VQerror=1mA ×100× 0.00099=99 μV

A mere 1% drift in the oscillator frequency generated a 0.056° phase shift in the analog filter.  This phase error forced the 100Ω resistance to bleed 99 µV of distortion into the Q-channel which is negligible to the target accuracy.

Excitation frequency: 100KHz

AHPF cutoff: 10KHz

Oscillator drift: 1% (1KHz)

Actual Resistance: 5000Ω

Excitation Current: 32μA

∆θ≈dθdfexc∆fexc=-1041052+1042103=-0.00099 radians=-0.056°

VQerror=32μA ×5000× 0.00099=158.4 μV

A mere 1% drift in the oscillator frequency generated a 0.056° phase shift in the analog filter.  This phase error forced the 5000 Ω resistance to bleed 158.4µV of distortion into the Q-channel which is not acceptable because if we divide this voltage error by your excitation current, it reports

Xerror=158.4µV 32μA=4.95Ω=5000× 0.00099

Which is significantly larger than 0.5Ω resolution.

Excitation frequency: 100KHz

AHPF cutoff: 10KHz

Oscillator drift: 1% (1KHz)

Actual Resistance: 100000Ω

Excitation Current: 32μA

∆θ≈dθdfexc∆fexc=-1041052+1042103=-0.00099 radians=-0.056°

VQerror=32μA ×100000× 0.00099=3.168 mV

A mere 1% drift in the oscillator frequency generated a 0.056° phase shift in the analog filter.  This phase error forced the 100000Ω resistance to bleed 3.168 mV of distortion into the Q-channel which is not acceptable because if we divide this voltage error by your excitation current, it reports

Xerror=3.168mV 32μA=99Ω

which is significantly larger than 0.5Ω resolution.

These results show that although oscillator drift has a negligible effect at low impedances, it produces substantial phase-induced measurement errors for large resistive loads. To maintain high accuracy over the full measurement range, the system should periodically recalibrate the analog phase response. One practical solution is to implement continuous background calibration by periodically switching the excitation path from the patient to a precision reference resistor for a short calibration interval.

# CMRR

The instrumentation amplifier must reject any common-mode voltage appearing simultaneously on both input electrodes. If the Common-Mode Rejection Ratio (CMRR) is insufficient, part of the common-mode signal is converted into an undesired differential voltage, directly degrading the impedance measurement accuracy.

The differential error produced by the common-mode voltage is

Verr=VcmCMRR

To ensure that this error does not obscure the smallest measurable impedance variation, the error voltage must remain smaller than the voltage corresponding to the minimum detectable impedance change:

Verr<Iexc∆Zmin

Substituting the expression for Verr gives

CMRR>VcmIexc∆Zmin

This equation defines the minimum CMRR required to ensure that common-mode interference remains below the system resolution. Assuming a common-mode voltage of 100 mV peak, the required CMRR can be calculated for different measurement conditions.

For a 100 kΩ body impedance, the excitation current is 10 µArms, and the minimum detectable impedance variation is 2.53 Ω. Therefore,

CMRR>0.1×√210μ 2.53=5589.777=74.95 dB

For a 25 Ω body impedance, the excitation current increases to 1.28 mArms, while the repeatability improves to 0.8 mΩrms. The required CMRR becomes

CMRR>0.1×√21.28m 0.8m=138106.79=102.8 dB

These calculations show that the low-impedance measurement presents the most demanding design condition because the smaller measurable impedance variation requires greater suppression of common-mode interference. Consequently, the INA should achieve a CMRR of approximately 103 dB or greater over the operating frequency range to ensure that common-mode-induced errors remain below the repeatability limit.

# Programmable Digital LPF and HPF Specifications

After analog IQ demodulation and analog low-pass filtering, the baseband I and Q signals are digitized by the sigma-delta ADC and further processed using programmable digital filters. These filters provide additional noise reduction and baseline stabilization without affecting the analog front-end.

The digital low-pass filter (DLPF) limits the measurement bandwidth, suppresses residual broadband noise, and improves measurement repeatability. Selecting a lower cutoff frequency reduces the integrated noise at the expense of a slower system response.

The digital high-pass filter (DHPF) removes the remaining DC component, slow baseline drift, and residual offset originating from the electrodes, analog circuitry, and IQ demodulation process. Unlike the analog HPF, which prevents saturation of the analog gain stages by rejecting large electrode offsets before amplification, the digital HPF operates on the digitized signal to stabilize the measurement baseline and improve long-term accuracy.

The cutoff frequencies of both filters are programmable through register settings, allowing the measurement bandwidth to be optimized for different bioimpedance applications. The available filter options are summarized in the following tables.

Table 5 Digital Low Pass Filter Cutoff Frequency (Hz)

| Register | Cutoff Frequency |
| --- | --- |
| 000 | Bypass |
| 001 | 0.005 × Sample Rate |
| 010 | 0.02 × Sample Rate |
| 011 | 0.08 × Sample Rate |
| 100 | 0.25 × Sample Rate |

Table 6 Digital High Pass Filter Cutoff Frequency (Hz)

| Register | Cutoff Frequency |
| --- | --- |
| 000 | Bypass |
| 001 | 0.00025 × Sample Rate |
| 010 | 0.002 × Sample Rate |

# Excitation Signal Specifications

The bioimpedance measurement system provides programmable excitation sources that can operate in either voltage or current stimulation mode. The excitation level is selected according to the expected body impedance to maximize the measurement signal-to-noise ratio while ensuring the input voltage remains within the allowable operating range. The available excitation settings are listed in the following tables.

Table 7 Stimulus Voltage Options

| Register | Voltage Magnitude (mVp) | Voltage Magnitude (mVrms) |
| --- | --- | --- |
| 00 | 50 | 35.4 |
| 01 | 100 | 70.7 |
| 10 | 250 | 177 |
| 11 | 500 | 354 |

The excitation current is selected according to the expected body impedance. Lower currents are used for high-impedance measurements to prevent excessive voltage, while higher currents improve the signal-to-noise ratio for low-impedance measurements.

Table 8 Stimulus Current Options

| Register | Stimulus Current |
| --- | --- |
| 0000 | 16nA rms |
| 0001 | 32nA rms |
| 0010 | 80nA rms |
| 0011 | 160nA rms |
| 0100 | 320nA rms |
| 0101 | 640nA rms |
| 0110 | 1.6uA rms |
| 0111 | 3.2uA rms |
| 1000 | 6.4uA rms |
| 1001 | 12.8uA rms |
| 1010 | 32uA rms |
| 1011 | 64uA rms |
| 1100 | 128uA rms |
| 1101 | 256uA rms |
| 1110 | 640uA rms |
| 1111 | 1.28mA rms |

In an ideal system, the excitation current is a pure sinusoid with constant amplitude. In practice, however, imperfections in the transmitter introduce small current fluctuations that appear as measurement errors after demodulation. These fluctuations can be divided into two categories depending on how they propagate through the receiver.

## Unmodulated Baseband Current Noise

The first category consists of low-frequency current fluctuations generated by the current mirrors, bias circuits and reference circuitry. This includes both thermal noise and flicker (1/f) noise that exist around DC. This noise is denoted by inlowt. Since it is not modulated onto the excitation carrier, it appears directly at low frequencies. After passing through the body impedance, it produces a low-frequency voltage

Vlowt=ZLinlowt

The Analog High-Pass Filter (AHPF) removes this component before it reaches the mixer. Consequently, the unmodulated baseband current noise has a negligible contribution to the final impedance measurement.

## Amplitude-Modulated Current Noise

The second category consists of noise that modulates the amplitude of the excitation carrier. Unlike the previous case, this noise rides on the carrier itself and therefore occupies narrow sidebands around the excitation frequency.

The total amplitude-modulated current noise is represented by inAMt. It is important to note that inAMt is not the DAC quantization noise alone. Instead, it represents the combined amplitude noise produced by the entire excitation path, including

- Current mirror thermal noise

- Current mirror flicker noise translated onto the carrier

- Reference voltage noise

- DAC quantization noise

- DAC glitch energy

- Clock jitter

- Finite DAC settling errors

The total excitation current can therefore be written as

Itotalt=I0+inAMtcosωct+inlowt

where I0 is the ideal excitation current amplitude. The excitation current passes through the body impedance ZL, generating the electrode voltage

Vint=ZLI0+inAMtcosωct+ZLinlowt

As discussed previously, the AHPF suppresses the low-frequency component ZLinlowt. Consequently, the signal entering the instrumentation amplifier is

VRXt=AINAZLI0+inAMtcosωct+θ0

The mixer multiplies this signal by cosωct, hence

Vmixt=AINAZLI0+inAMtcosωct+θ0cosωct

After low-pass filtering:

Vbbt=AINAZL2I0+inAMtcosθ0

Expanding the expression:

Vbbt=AINAZL2I0cosθ0+AINAZL2inAMtcosθ0

The first term is the desired demodulated impedance signal. The second term is the error introduced by amplitude fluctuations of the excitation current. Since these fluctuations are modulated onto the carrier, they pass through the AHPF together with the desired signal and are translated directly to baseband by the synchronous mixer. Consequently, they cannot be removed by the receiver filtering.

## DAC Quantization Noise

One contributor to the total amplitude-modulated current noise is the finite resolution of the excitation DAC.

For an ideal 10-bit DAC,

iLSB=IFS210

where IFS is the full-scale excitation current. The corresponding RMS quantization noise is

iqrms=iLSB12

This noise is approximately white and uniformly distributed over the Nyquist bandwidth (Fs/2) where Fs is the DAC sampling rate. Therefore, the equivalent current-noise spectral density is

eqDAC=iqrmsFs/2

Only a very small portion of this broadband noise falls within the physiological measurement bandwidth after synchronous demodulation and digital filtering. The integrated in-band quantization noise is therefore

iqfinal=eqDACBW

where BW is the equivalent baseband noise bandwidth (approximately 4 Hz). Finally, this current noise produces an equivalent impedance error

∆ZDAC=ZLiqfinalIexc

This equation shows that the impedance error caused by DAC quantization noise increases with the measured body impedance and decreases with excitation current. It also demonstrates why the highest impedance measurement range places the most stringent requirements on the excitation source, whereas the lower impedance ranges are considerably more tolerant of DAC noise.

Table 9 Impedance error caused by DAC for different Iexc and ZL=25Ω

| Iexc (rms) | Iref | iLSB | iqrms | Fs | eqDAC | BW | iqfinal | ZL | ∆ZDAC |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 16nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 312 µΩ |
| 32nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 156 µΩ |
| 80nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 62.3 µΩ |
| 160nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 31 µΩ |
| 320nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 15.6 µΩ |
| 640nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 7.79 µΩ |
| 1.6uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 3.1 µΩ |
| 3.2uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 1.56 µΩ |
| 6.4uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 0.78 µΩ |
| 12.8uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 0.39 µΩ |
| 32uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 0.16 µΩ |
| 64uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 77.9 nΩ |
| 128uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 38.9 nΩ |
| 256uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 1.95 nΩ |
| 640uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 7.8 nΩ |
| 1.28mA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 25Ω | 3.9 nΩ |

Table 10 Impedance error caused by DAC for different Iexc and ZL=5kΩ

| Iexc (rms) | Iref | iLSB | iqrms | Fs | eqDAC | BW | iqfinal | ZL | ∆ZDAC |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 16nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 0.063 Ω |
| 32nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 0.03Ω |
| 80nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 0.013Ω |
| 160nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 6.3 mΩ |
| 320nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 3.1 mΩ |
| 640nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 1.56 mΩ |
| 1.6uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 0.63 mΩ |
| 3.2uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 0.31 mΩ |
| 6.4uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 0.16 mΩ |
| 12.8uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 78 µΩ |
| 32uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 31.2 µΩ |
| 64uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 16 µΩ |
| 128uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 7.8 µΩ |
| 256uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 3.9 µΩ |
| 640uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 1.6 µΩ |
| 1.28mA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 5kΩ | 0.78 µΩ |

Table 11 Impedance error caused by DAC for different Iexc and ZL=50kΩ

| Iexc (rms) | Iref | iLSB | iqrms | Fs | eqDAC | BW | iqfinal | ZL | ∆ZDAC |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 16nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 0.63 Ω |
| 32nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 0.31 Ω |
| 80nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 0.12 Ω |
| 160nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 0.063 Ω |
| 320nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 0.031 Ω |
| 640nA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 0.016 Ω |
| 1.6uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 6 mΩ |
| 3.2uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 3 mΩ |
| 6.4uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 1.6 mΩ |
| 12.8uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 779 µΩ |
| 32uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 312 µΩ |
| 64uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 156 µΩ |
| 128uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 78 µΩ |
| 256uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 39 µΩ |
| 640uA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 16 µΩ |
| 1.28mA | 250nA | 0.244nA | 70.47pA | 1MHz | 99.67f | 4Hz | 0.199pA | 50kΩ | 7.8 µΩ |

# Analog HPF and Programmable Gains

The analog front-end consists of a programmable high-pass filter (HPF), an instrumentation amplifier (INA), and a programmable gain amplifier (PGA). The HPF removes the large DC electrode offset and low-frequency baseline variations while allowing the high-frequency excitation carrier to pass with minimal attenuation. The INA provides high common-mode rejection and amplifies the small differential bioimpedance signal. Finally, the PGA provides programmable gain so that signals corresponding to different body impedances fully utilize the ADC input range without causing saturation.

The HPF cutoff should be selected well below the excitation frequency to avoid attenuating the carrier while effectively rejecting electrode DC offsets and motion-induced baseline drift. The INA and PGA gains are then programmed according to the expected body impedance to maximize the signal presented to the sigma-delta ADC while maintaining sufficient headroom for offsets and noise.

Table 12 Analog High Pass Filter

| Register | Cutoff frequency |
| --- | --- |
| 000 | 100Hz |
| 001 | 200Hz |
| 010 | 500Hz |
| 011 | 1kHz |
| 100 | 2kHz |
| 101 | 5kHz |
| 110 | 10kHz |

Table 13 INA and PGA Gain Settings

| Register | Gain |
| --- | --- |
| 00 | 1 |
| 01 | 2 |
| 10 | 5 |
| 11 | 10 |

# Guard Amplifier

The core principle of a guard amplifier is active shielding, which neutralizes the effective capacitance by ensuring zero potential difference across the parasitic capacitor. Imagine a single electrode connected to a receive electrode.

Let Vsig be the actual voltage at the electrode-tissue interface.

Let Rs be the source impedance (electrode-tissue contact resistance).

Let Cp be the parasitic capacitance of the PCB trace to ground.

Let Vin be the voltage arriving at the input of the INA.

This setup forms a standard RC LPF, therefore:

Vin=Vsig11+RsCp

We can find the phase lag introduced by the parasitic capacitance:

θ=-tan-1ωRsCp

If the contact resistances are mismatched or the parasitic capacitances are mismatched, the phase lag on the positive and negative input will be different:

∆θ=tan-1ωRsNCpN-tan-1ωRsPCpP

The current Ic flowing from the signal trace into the parasitic capacitance

Ic=Vin-VguardZcap=

Since the guard amplifier drives the guard trace with the input voltage,

Vguard=AVin

Therefore,

Ic=Vin-VguardZcap=Vin(1-A)SCsg

Now we can define the effective capacitance seen by the signal trace:

Ceff=(1-A)Csg

If the guard amplifier is an ideal unity-gain buffer, the gain A=1, thus

Ceff=1-ACsg=0

Because the effective parasitic capacitance is driven to zero, the transfer function of the input node becomes:

Vin=Vsig11+RsCp≈Vsig

θ=-tan-1ωRsCp=0°

To put it simply, by actively driving the guard trace to equal the signal trace voltage, the potential difference across the trace-to-trace dielectric is forced to zero. According to Q=C∆V, if ∆V=0, no charge Q is transferred. Mathematically, the pole is pushed to an infinitely high frequency (or practically, far outside the bandwidth of interest), nullifying the phase error regardless of how poorly matched the electrode contact resistances (Rs) are.

# Bioimpedance Applications

Bioimpedance (BioZ) measures the electrical impedance of biological tissue by injecting a small alternating current and measuring the resulting voltage. Because different tissues exhibit different electrical properties and these properties vary with physiological activity, bioimpedance has become a versatile sensing technique in medical devices and wearable electronics.

Typical BioZ applications include:

- Body Composition Analysis (BIA)

- Impedance Cardiography (ICG)

- Respiration Monitoring (Pneumography)

- Galvanic Skin Response (GSR) / Electrodermal Activity (EDA)

- Hydration Assessment

- Continuous physiological monitoring in wearable devices, such as smartwatches.

The required electrode configuration, excitation frequency, current amplitude, and measurement duration depend on the intended application.

## Four-Electrode Configuration

The four-electrode (tetrapolar) configuration is the preferred method for accurate bioimpedance measurements because it separates the current injection electrodes from the voltage sensing electrodes. Two electrodes inject a small AC current into the body, while a second pair measures the resulting voltage drop across the tissue. Since virtually no current flows through the voltage-sensing electrodes, the measured impedance is largely independent of the electrode contact resistance and skin-electrode interface impedance. This significantly improves measurement accuracy, particularly for body composition and hydration analysis.

The four-electrode configuration is widely used in

- Body Composition Analysis (BIA)

- Hydration Assessment

- Bioimpedance Spectroscopy (BIS)

- High-accuracy wearable bioimpedance devices

In wearable devices such as smartwatches, placing four conventional electrodes on the wrist is impractical because only one side of the body is in contact with the watch. Instead, modern smartwatches implement the tetrapolar measurement by combining electrodes on the underside of the watch with conductive metal electrodes integrated into the watch bezel or side buttons. The bezel electrode is therefore not a special electrode type. It simply refers to the conductive metal frame, bezel, or buttons of the smartwatch that are reused as bioimpedance electrodes.

During a body composition measurement, the user wears the watch on one wrist while placing two fingers of the opposite hand on the metallic bezel or side buttons. This creates a hand-to-hand electrical path through the upper body. Typically,

- two electrodes located on the underside of the watch contact the wrist,

- two conductive electrodes integrated into the bezel or side buttons contact two fingers of the opposite hand.

Together these four contact points form a complete tetrapolar bioimpedance measurement system. One electrode pair injects the excitation current, while the second pair senses the resulting voltage. Because the voltage measurement is isolated from the current injection path, the influence of contact impedance is greatly reduced compared with a two-electrode measurement.

This architecture is used in several commercial smartwatches for estimating physiological parameters such as

- Total Body Water (TBW),

- Fat-Free Mass (FFM),

- Body Fat Percentage (BF%), and

- Skeletal Muscle Mass.

A recent study reported a commercial smartwatch using two electrodes underneath the watch and two electrodes integrated into the watch frame to perform single-frequency tetrapolar bioimpedance measurements at 50 kHz, enabling hand-to-hand body composition estimation through proprietary algorithms.

Figure 1 4-Electrode configuration in a smartwatch

## Why is 50 kHz the Standard Excitation Frequency?

The electrical behaviour of biological tissue depends strongly on frequency because cell membranes behave as capacitors.

At very low frequencies (1-5 kHz), the membrane capacitance exhibits a large reactance, preventing current from entering the cells. The excitation current therefore flows almost entirely through the extracellular fluid. Consequently, low frequencies primarily measure ExtraCellular Water (ECW).

At high frequencies (>200 kHz), the membrane reactance becomes very small and current passes through both intracellular and extracellular fluids. In this case the measurement approaches the total body conductivity, and the contribution of membrane capacitance becomes less distinguishable.

An excitation frequency of approximately 50 kHz provides an excellent compromise.

At this frequency,

- sufficient current penetrates the cells,

- cell membranes still produce measurable phase shift,

- both resistance and reactance can be measured accurately.

Therefore, Resistance (R) mainly reflects conductive body fluids and is used to estimate Total Body Water (TBW) and Fat-Free Mass (FFM) while Reactance (Xc) originates primarily from cell membrane capacitance and provides information related to body cell mass, and cellular health.

For these reasons, 50 kHz has become the de facto standard excitation frequency for bioimpedance analysers.

## Impedance Cardiography (ICG)

Impedance Cardiography measures the small impedance variations produced by blood flowing through the thoracic cavity.

Because blood has significantly higher conductivity than surrounding tissue, the thoracic impedance changes during each heartbeat.

These impedance changes are used to estimate

- Stroke Volume (SV)

- Cardiac Output (CO)

- Cardiac Contractility

- Thoracic Fluid Content

Unlike body composition measurements, ICG is performed continuously. The physiological signal occupies approximately 0.05–65 Hz which contains heart rate and blood flow waveform.

## Respiration Monitoring

This is the tracking of the breathing cycle (respiration rate and depth) by monitoring the physical expansion and contraction of the lungs. Similar to ICG, this measurement relies on thoracic bioimpedance. Air is a poor electrical conductor. When a patient inhales and the lungs fill with air, the electrical impedance across the chest increases. When they exhale, the impedance decreases. Because breathing is a much slower physical process than a heartbeat, the impedance changes occur at very slow frequencies (typically under 2 Hz) but have a larger amplitude than the cardiovascular impedance ripples.

## GSR/EDA Measurement

Galvanic Skin Response (GSR), also known as Electrodermal Activity (EDA), measures changes in skin conductance caused by sweat gland activity controlled by the sympathetic nervous system.	Unlike body composition measurements, GSR is primarily interested in changes in skin resistance rather than absolute body impedance. Therefore, a two-electrode configuration is generally sufficient. The measured impedance naturally includes

- skin impedance,

- electrode contact impedance,

- tissue impedance.

Because GSR algorithms monitor relative changes over time, the contact impedance is usually acceptable and does not significantly reduce measurement performance.

Typical applications include

- stress monitoring,

- emotion recognition,

- sleep analysis.

## Continuous and Discontinuous Bioimpedance Measurements

Bioimpedance applications can also be classified according to whether the impedance is measured once or continuously over time.

Discontinuous measurements are performed only when the user initiates a measurement. The objective is to estimate a static physiological parameter from a single impedance snapshot. Typical examples include body composition analysis and hydration assessment. These measurements usually require only a few seconds and are repeated infrequently.

In contrast, continuous measurements monitor impedance changes in real time to capture physiological activity. The excitation current is applied continuously, and impedance is sampled repeatedly over an extended period. This enables slow or rapid physiological variations to be tracked.

Typical continuous applications include:

- Impedance Cardiography (ICG): continuous monitoring of cardiac-induced impedance changes for estimating stroke volume and cardiac output.

- Respiration Monitoring: continuous measurement of thoracic impedance variations caused by inhalation and exhalation.

- Galvanic Skin Response (GSR)/Electrodermal Activity (EDA): continuous monitoring of changes in skin conductance associated with sympathetic nervous system activity.

# BMS3 Limitations (Digital DFT architecture)

The external connections for the impedance measurement system are shown above. Here, R0 is the output resistance contributing from Op Amp, switches, and pads. The system can measure magnitude of impedance (|Z|) from 1 kΩ to 300 kΩ impedance and phase from -90 deg to +90 deg.

The output voltage of the current-to-voltage converter (I-to-V) at RFB pin for unknown impedance and calibration resistor selection can be represented as:

VRFB=RFBZmeasVout

VRFB=RFBRcalVout

Resolution is the smallest physical change in the input Zmeas that results in a detectable change in the output (a bit flip in the ADC). Resolution is derived from the system’s sensitivity, which is the derivative of the transfer function. Thus, the sensitivity of the system mathematically defined as the derivate of the transfer function with respect to Zmeas.

dVRFBdZmeas=-RFBZmeas2Vout

The slop of the response dictates the measurement accuracy across different impedance ranges. The slope is much steeper at the lower end of the Zmeas range compared to the higher end.

To find the absolute impedance resolution ∆Zmeas, we relate this derivative to the quantization limit of the ADC.

VLSB=VFS2N

By substituting the minimum resolvable voltage change (∆VRFB=VLSB) into derivate equation and solving for the impedance change, we obtain the absolute resolution limit:

∆Zmeas=VLSBZmeas2VoutRFB

The system resolution is better (smaller ∆Zmeas) for smaller impedances. The absolute error or resolution limit degrades proportionally to the square of the unknown impedance (Zmeas2). As the measured impedance increases, a much larger change in impedance is required to produce a VLSB change at the ADC input.

Relative Error defines how far the final calculated value deviates from absolute reality, expressed as a fraction or percentage.

Relative Error=E=∆ZmeasZmeas=VLSBZmeasVoutRFB

Zmeas=EVoutRFBVLSB

This confirms that the relative error increases linearly as Zmeas becomes larger than RFB.

Now the question is that for what impedance range using 10-bit ADC 0.5% relative error is achievable for 900mV peak to peak excitation voltage and RFB=20 kΩ?

VLSB=1.5210=1.4648 mV

Zmeas=0.005450m× 20k1.4648 m=30.7209 kΩ

However, an ENOB of 10-bit ADC can’t reach 10 and it is typically 1 or 2 bits lower. For ENOB of 9 bits, Zmeas from 30.7 kΩ drops to approximately 15 kΩ.

The lower limit of the range is dictated by the saturation limit of the ADC. The ADC is biased at VDD/2=0.75V, therefore, can’t measure a signal swinging more than ±0.75V without clipping. Using the transfer function to find the minimum impedance before clipping happens:

VRFB=RFBZmeasVout≤0.75

Zmeas≥450m× 20k0.75=12 kΩ

For RFB=10kΩ, the impedance range is

Zmeas≤0.005450m× 10k1.4648 m=15.36045 kΩ

Zmeas≥450m× 10k0.75=6 kΩ

The impedance measurement range is limited to

ZmeasMAX=EmaxVoutRFBVLSB

ZmeasMIN=VoutRFBVDD2

The PGA after TIA only shifts both ZmeasMAX and ZmeasMIN by its gain (G) as follows:

ZmeasMAX=EmaxVoutRFBVLSBG

ZmeasMIN=GVoutRFBVDD2

Table 14 summarizes the impedance range for different values of ENOB and PGA gain where VFS=VDD=1.5V. As table 1 shows, for RFB of 20 kΩ, only for the impedance range of 12000 to 30720 Ω relative error of 0.005 (=0.5%) is achievable. If gain is changed from 1 to 2, the range shifts as discussed above. If you change the Vout from 0.9/2 to 0.1/2 V, as there are two options, the impedance range changes from 2000 to 5120 Ω. But in projects before, the gain option is only 1 and 5 which means it is not possible to measure a wide range of impedance with relative error of 0.005 when gain is changed from 1 to 5. This is the first issue that the designer (Shamsul) did not realize. The second issue is in the range of impedance measurements for 0.5% accuracy. The designer reported a range of 20000 to 40000 Ω.

Table 14 Impedance range and corresponding resolution for 10-bit ENOB, RFB=20 kΩ, and E=0.5%

| ENOB (bit) | VLSB(V) | E | Vout(V) | G | ZMAX(Ω) | ZMIN(Ω) | ResZMAX(Ω) | ResZMIN(Ω) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 10 | 0.001465 | 0.005 | 0.45 | 1 | 30720 | 12000 | 153.6 | 23.4375 |
| 10 | 0.001465 | 0.005 | 0.45 | 2 | 61440 | 24000 | 614.4 | 93.75 |
| 10 | 0.001465 | 0.005 | 0.45 | 3 | 92160 | 36000 | 1382.4 | 210.9375 |
| 10 | 0.001465 | 0.005 | 0.45 | 4 | 122880 | 48000 | 2457.6 | 375 |
| 10 | 0.001465 | 0.005 | 0.45 | 5 | 153600 | 60000 | 3840 | 585.9375 |
| 10 | 0.001465 | 0.005 | 0.075 | 1 | 5120 | 2000 | 25.6 | 3.90625 |
| 10 | 0.001465 | 0.005 | 0.075 | 2 | 10240 | 4000 | 102.4 | 15.625 |
| 10 | 0.001465 | 0.005 | 0.075 | 3 | 15360 | 6000 | 230.4 | 35.15625 |
| 10 | 0.001465 | 0.005 | 0.075 | 4 | 20480 | 8000 | 409.6 | 62.5 |
| 10 | 0.001465 | 0.005 | 0.075 | 5 | 25600 | 10000 | 640 | 97.65625 |

Table 15 summarizes the impedance range for the same parameters, but ENOS is assumed 8 because a conventional 10-bit SAR ADC ENOB in typical corner is 9 and in extreme cases may reach 8 (slow-slow or fast-fast corner). As Table 15 shows, for RFB of 20 kΩ, system can’t achieve 0.5% relative accuracy as ZMAX < ZMIN. For ENOB of 8 bits, Table 16 shows the range of impedances with 1% accuracy.

Table 15 Impedance range and corresponding resolution for 8-bit ENOB, RFB=20 kΩ and E=0.5%

| ENOB (bit) | VLSB(V) | E | Vout(V) | G | ZMAX(Ω) | ZMIN(Ω) | ResZMAX(Ω) | ResZMIN(Ω) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 8 | 0.00586 | 0.005 | 0.45 | 1 | 7680 | 12000 | 38.4 | 93.75 |
| 8 | 0.00586 | 0.005 | 0.45 | 2 | 15360 | 24000 | 153.6 | 375 |
| 8 | 0.00586 | 0.005 | 0.45 | 3 | 23040 | 36000 | 345.6 | 843.75 |
| 8 | 0.00586 | 0.005 | 0.45 | 4 | 30720 | 48000 | 614.4 | 1500 |
| 8 | 0.00586 | 0.005 | 0.45 | 5 | 38400 | 60000 | 960 | 2343.75 |
| 8 | 0.00586 | 0.005 | 0.075 | 1 | 1280 | 2000 | 6.4 | 15.625 |
| 8 | 0.00586 | 0.005 | 0.075 | 2 | 2560 | 4000 | 25.6 | 62.5 |
| 8 | 0.00586 | 0.005 | 0.075 | 3 | 3840 | 6000 | 57.6 | 140.625 |
| 8 | 0.00586 | 0.005 | 0.075 | 4 | 5120 | 8000 | 102.4 | 250 |
| 8 | 0.00586 | 0.005 | 0.075 | 5 | 6400 | 10000 | 160 | 390.625 |

Table 16 Impedance range and corresponding resolution for 8-bit ENOB, RFB=20 kΩ, and E=1%

| ENOB (bit) | VLSB(V) | E | Vout(V) | G | ZMAX(Ω) | ZMIN(Ω) | ResZMAX(Ω) | ResZMIN(Ω) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 8 | 0.00586 | 0.01 | 0.45 | 1 | 15360 | 12000 | 153.6 | 93.75 |
| 8 | 0.00586 | 0.01 | 0.45 | 2 | 30720 | 24000 | 614.4 | 375 |
| 8 | 0.00586 | 0.01 | 0.45 | 3 | 46080 | 36000 | 1382.4 | 843.75 |
| 8 | 0.00586 | 0.01 | 0.45 | 4 | 61440 | 48000 | 2457.6 | 1500 |
| 8 | 0.00586 | 0.01 | 0.45 | 5 | 76800 | 60000 | 3840 | 2343.75 |
| 8 | 0.00586 | 0.01 | 0.075 | 1 | 5120 | 2000 | 51.2 | 7.8125 |
| 8 | 0.00586 | 0.01 | 0.075 | 2 | 10240 | 4000 | 204.8 | 31.25 |
| 8 | 0.00586 | 0.01 | 0.075 | 3 | 15360 | 6000 | 460.8 | 70.3125 |
| 8 | 0.00586 | 0.01 | 0.075 | 4 | 20480 | 8000 | 819.2 | 125 |
| 8 | 0.00586 | 0.01 | 0.075 | 5 | 25600 | 10000 | 1280 | 195.3125 |

The designer suggested using 0.15V peak to peak excitation voltage to measure the impedance in the range of 1 kΩ to 20 kΩ. Table 11 shows the impedance range for RFB=1000 Ω , ENOB of 10 bits. For RFB=20 kΩ, we can’t find a range even for 1% accuracy.

Table 17 Impedance range and corresponding resolution for 8-bit ENOB, E=1%, RFB=1000 Ω, and Zmeas < 10000.

| ENOB (bit) | VLSB(V) | E | Vout(V) | G | ZMAX(Ω) | ZMIN(Ω) | ResZMAX(Ω) | ResZMIN(Ω) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 10 | 0.001465 | 0.01 | 0.075 | 1 | 512 | 100 | 5.12 | 0.1953125 |
| 10 | 0.001465 | 0.01 | 0.075 | 2 | 1024 | 200 | 20.48 | 0.78125 |
| 10 | 0.001465 | 0.01 | 0.075 | 3 | 1536 | 300 | 46.08 | 1.7578125 |
| 10 | 0.001465 | 0.01 | 0.075 | 4 | 2048 | 400 | 81.92 | 3.125 |
| 10 | 0.001465 | 0.01 | 0.075 | 5 | 2560 | 500 | 128 | 4.8828125 |
| 8 | 0.00586 | 0.01 | 0.075 | 1 | 128 | 100 | 1.28 | 0.78125 |
| 8 | 0.00586 | 0.01 | 0.075 | 2 | 256 | 200 | 5.12 | 3.125 |
| 8 | 0.00586 | 0.01 | 0.075 | 3 | 384 | 300 | 11.52 | 7.03125 |
| 8 | 0.00586 | 0.01 | 0.075 | 4 | 512 | 400 | 20.48 | 12.5 |
| 8 | 0.00586 | 0.01 | 0.075 | 5 | 640 | 500 | 32 | 19.53125 |

If we target an impedance range of 25 Ω, we need to have two PGA capable of driving 3mA. For RFB=200 Ω, the range is limited to ZMAX=128 Ω.

Table 18 Impedance range and corresponding resolution for 8-bit ENOB, E=1%, RFB=1000 Ω, and Zmeas  < 10000.

| ENOB (bit) | VLSB(V) | E | Vout(V) | G | ZMAX(Ω) | ZMIN(Ω) | ResZMAX(Ω) | ResZMIN(Ω) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 8 | 0.00586 | 0.01 | 0.075 | 1 | 25.6 | 20 | 0.256 | 0.15625 |
| 8 | 0.00586 | 0.01 | 0.075 | 2 | 51.2 | 40 | 1.024 | 0.625 |
| 8 | 0.00586 | 0.01 | 0.075 | 3 | 76.8 | 60 | 2.304 | 1.40625 |
| 8 | 0.00586 | 0.01 | 0.075 | 4 | 102.4 | 80 | 4.096 | 2.5 |
| 8 | 0.00586 | 0.01 | 0.075 | 5 | 128 | 100 | 6.4 | 3.90625 |

It is worth noting that the DAC output resistance (R0) limits the impedance range. For impedances lower than R0 the voltage mostly drops across R0 rather than the device under test ( Rcal or Zmeas).  However, the designer claimed that the system can measure low impedances such as 1kΩ with R0=780 Ω (from his report) with 0.5% accuracy.

VRFB=10001780×0.15=0.084Vp2p

If we divide 84mV peak to VLSB=0.001465 V, the signal spans about 57 or 58 VLSB out of the available 1024.

## Phase Measurements

Aperture jitter of the sampling clock constraints the accuracy and should not introduce a voltage error greater than half of VLSB. Assuming the excitation signal is a pure sine wave, the voltage at the ADC input is:

Vint=Vmsin2πfint

The maximum slew rate (the point where the jitter causes the largest voltage error) occurs when

dVintdt=2πfinVmcos2πfint=2πfinVm

To ensure the voltage error remains below, the jitter time must satisfy the following inequality:

tj<0.5×VLSB2πfinVm=0.52πfinVmVFS2N

If signal amplitude is parameterized as k-times the half full-scale voltage

Vm=kVFS2

where k<1, then

tj<0.52πfinVmVFS2N=0.52πfinkVFS2VFS2N=12πfink12N

The accuracy of this phase measurement relies heavily on the strict timing synchronization between the transmitter DAC and receiver ADC stages. To ensure correct phase measurement (which is nonzero for RC unknown impedances), the ADC enable must occur at an exact, fixed time after the DAC is activated. Because the signal generation and the ADC sampling are both driven by the master clock, any shift in the clock frequency skews this critical timing window.

The system generates the output signal frequency based on the master clock and a user-selected DAC multiplier (M). Thus

∆fsig=∆fclk4×M

The system is programmed to wait a specific settling time between enabling the DAC and enabling the ADC to ensure accurate sampling.  This settling time is defined as a Number of Cycles (NOC) of the output signal. However, because the clock has drifted, the actual time that passes during this delay is governed by the actual output signal frequency (fsig'), which is driven by the actual drifted clock frequency. Therefore,

fclk'=fclk±∆fclk

First, we define the actual signal frequency:

fsig'=fclk±∆fclk4×M

Next, we calculate the absolute time delay it takes to complete the programmed NOC:

tdelay=NOCfsig'

By substituting fsig' into the time delay equation, we get:

tdelay=NOC×4×Mfclk±∆fclk

Now we substitute the expanded formulas for both the frequency error and the actual time delay back into phase error introduced by the clock jitter:

θDFT=360°×∆fsig×tdelay=360°×∆fclk4×M×NOC×4×Mfclk±∆fclk=360°×∆fclkfclk±∆fclk×NOC

This equation demonstrates that the phase error caused by clock shift is independent of the chosen signal frequency and is strictly a product of the clock error and the length of the settling time.

When measuring complex RC series impedances, the physical physics of the load can break the measurement if the excitation frequency is too high. The time constant of the series impedance must be lower than the period of the applied signal frequency. If the time constant is too high, the voltage across the capacitor can’t follow the input sinusoidal signal. To prevent this failure, the system enforces a strict frequency limit constraint:

∆fsig≤12πRC

In an ideal op amp, the bandwidth is infinite, and a signal passing through a non-inverting configuration experiences exactly zero degrees of phase shift. In reality, a limited gain bandwidth (GBW) forces the amplifier to act as a LPF, introducing a frequency dependent phase lag.

To Ensure stability, most op amp are internally compensated with a dominant pole. This causes the open-loop gain AOL, to roll off at -20 dB/dec. The open loop response is modelled mathematically as a single-pole LPF:

AOLf=A01+jffb

where A0 is the DC open-loop gain, fb is the open loop dominant pole (-3 dB frequency), and f is the operating signal frequency. GBW product is defined as

GBW=fbA0

When we apply negative feedback to the amplifier using a feedback factor β, we calculate the closed-loop gain using control theory:

ACLf=AOLf1+βAOLf

By substituting the open loop model:

ACLf=A01+jffb1+βA01+jffb=A01+jffb+βA0=A01+βA0+jffb=A01+βA01+jffb1+βA0

This reveals that the ideal closed-loop DC gain (ACL0) is ≈1β and the closed loop pole is

fc=fb1+βA0

Since A0 is considerably large, 1+βA0≈βA0. Therefore,

fc≈fbβA0=β GBW=GBWACL0

Therefore

ACLf=A01+jffb1+βA01+jffb=A01+jffb+βA0=A01+βA0+jffb=ACL01+jffc

The phase shift introduced by the closed-loop system is:

∆φ=-tan-1ffc=-tan-1fACL0GBW

In high precision analogue design, phase error is very small

∆φrad≈-fACL0GBW

∆φdeg≈-fACL0GBW180π

The closed-loop gain and thus the phase accuracy is directly dependent on the ratio of RFB and Zmeas. Because:

β=ZmeasRFB+Zmeas

therefore,

ACL0=RFB+ZmeasZmeas=1+RFBZmeas

This is the formula for a non-inverting amplifier. In op amp terminology, this is universally called the Noise Gain (Anoise). It is called noise gain because the op amp’s own internal input voltage noise is always amplified by this non-inverting factor, regardless of how you apply signals.

Let assume RFB=20 kΩ and GBW of 10MHz and phase error of less than 0.1°. Using the

tan-1fACL0GBW≤0.1°

fACL0GBW≤tan0.1°

If GBW is 1MHz then

fACL0106≤tan0.1°=0.0017453

ACL0≤1745.3f

20000Zmeas≤1745.3f-1

200001745.3f-1≤Zmeas

Thus

1745.3f-1>0

f<1745.3

Therefore, with a 1MHz GBW op amp, it is physically impossible to achieve a 0.1° phase accuracy for any signal frequency higher than 1.74 kHz. At f=0.9766 kHz:

200001745.3976.6-1=25409≤Zmeas

Thus, to maintain 0.1° accuracy at f=0.9766 kHz, Zmeas must be strictly greater than 25.4 kΩ. If GBW= 10MHz,

fACL0107≤tan0.1°=0.0017453

ACL0≤17453f

20000Zmeas≤17453f-1

2000017453f-1≤Zmeas

For Zmeas to be positive

17453f-1>0

f<17453

With a 10 MHz GBW, the op amp can now support a 0.1° phase accuracy for signals up to 15.45 kHz. At 3.9063 kHz

20000174533906.3-1=5767≤Zmeas

At 15.625 kHz

200001745315625-1=170940≤Zmeas

This phase error caused by a finite GBW product is applicable to the TX (DAC) PGA, TIA, and RX PGA as well. In fact, it applies to every operational amplifier operating in a closed loop gain within the signal chain. Because all these amplifiers operate in series, the total phase error is the superposition of all individual phase shifts:

∆φrad=∆φDACPGA+∆φTIA+∆φRXPGA

The DAC PGA acts as a voltage amplifier to drive the excitation signal into the load. If the PGA’s gain is altered to switch between output voltage ranges, its closed-loop bandwidth changes, which means the phase lag shifts proportionally. Based on the previous derivation:

∆φDACPGA=≈-fADACPGA_GAINGBWDACPGA

If you achieve a higher output drive by increasing the amplifier’s gain, the phase lag of the DAC PGA will increase because the closed-loop pole moves closer to the operating frequency.

Phase lag of the TIA changes dynamically depending on the magnitude of the impedance you are measuring.  Because Zmeas is a variable in your system, the TIA’s phase error is not constant.

Phase error contributed by the RX PGA is

∆φRXPGA=≈-fARXPGA_GAINGBWRXPGA

It is worth noting that Rcal does not calibrate this phase error for reactive loads. Swapping a purely resistive Rcal for a reactive Zmeas alters the noise gain of the TIA, shifting its closed-loop pole and changing the phase lag dynamically.

To measure a 30nF capacitor (largest asked by the customer) with a 20 kΩ feedback resistor, we must calculate the noise waregain for each frequency reported in the document (0.9766, 1.953, 3.9063, 7.8125, 15.625, 31.25 kHz).

ACL0=1+RFBZmeas=1+2πfRFBC=1+0.00377f

fACL00.0017453≤GBW

For f= 0.9766 kHz:

ACL0=4.68

2618740 Hz≤GBW

Table 14 summarizes the ACL0 and GBW required for different excitation frequencies. This illustrates why driving large capacitive loads at high frequencies is so challenging in analogue design. By the time, frequency is 31.25 kHz, the 30nF capacitor impedance drops to 169.77 Ω. Because RFB is fixed at 20 kΩ, the TIA is forced to operate a massive noise gain of almost 118. To support such a gain a gain bandwidth of more than 2.12 GHz is needed. For a two-stage op amp, a tail current of 21.36 mA is required which is unfeasible for a standard precision measurement AFE.

Table 19 Required GBW required for TIA for different excitation frequencies for a 30nF capacitive load

| Frequency (Hz) | ACL0 | GBW(kHz) |
| --- | --- | --- |
| 976.6 | 4.681782 | 2619.737754 |
| 1953 | 8.36281 | 9358.028952 |
| 3906.3 | 15.726751 | 35199.33962 |
| 7812.5 | 30.453125 | 136317.5609 |
| 15625 | 59.90625 | 536317.6281 |
| 31250 | 118.8125 | 2127365.281 |
| 50000 | 189.5 | 5428866.098 |

Why IQ for BAF4P1

IQ demodulation advantages over DFT processing:

- In the IQ method, the high frequency voltage is multiplied (mixed) with the reference oscillator signals (In-phase and Quadrate) before the ADC. This mixing process down converts the excitation frequency (Fexc) to DC or a very low frequency (Zero-IF). Because the carrier frequency is removed, the ADC only needs to sample the physiological bandwidth of the impedance changes, which is typically very low. Since the required bandwidth is so low, you can utilize a highly oversampled sigma delta ADC. This allows you to achieve a higher ENOB with very low-power consumption. High ENOB is critical in BioZ because the small physiological impedance variations are riding on top of a massive static base impedance. In DFT method, ADC directly digitizes the high frequency AC BioZ signal. To accurately capture the signal, the ADC sampling rate must strictly obey the Nyquist theorem. At these higher sampling rates, using a high-resolution ADC becomes power-hungry and difficult to design. You are forced to use SAR ADC or a pipelined ADC. That should be designed from scratch; however, a sigma delta ADC is already designed for BAF4 project. High frequency sampling is much more susceptible to clock jitter, which degrades the SNR. Analog mixers in the IQ method introduce non-idealities such as IQ phase/gain mismatch which needs calibration. DFT processing avoids this issue.

Changes on the TX side

- BMS3 comes with only voltage excitation. Current excitation is the gold standard for clinical BioZ and dynamic tissue measurement. There is Strict Patient Safety (IEC 60601-1 compliance). When interacting with living tissue, safety standards strictly limit the maximum allowable alternating current that can pass through the body clamped microamps depending on the frequency. When applying a voltage or current to a human body, patient safety must be the top priority. According to IEC 60601-1, the maximum allowable AC current under normal condition is 100uA at low frequencies such as 50 or 60Hz.  Table below shows the allowable current magnitudes at the frequency range.

| Frequency Range (Hz) | Maximum Current (uA) |
| --- | --- |
| <512 | 64 |
| >512 | 128 |
| >2048 | 256 |
| >8192 | 640 |
| >16384 | 1280 |

If you apply a fixed voltage, the resulting current is entirely at the mercy of the load impedance. In human bodies, skin impedance is highly unstable. If the subject sweats, skin impedance drops exponentially, causing a massive potentially unsafe spike in the injected current. Therefore, a programmable current source should be implemented in BAF4P1.  Voltage excitation should yet be included.  Because current sources have a physical limit called compliance voltage. Suppose you are measuring a highly resistive dry electrochemical sensor with Z=5M ohm. If you attempt to push just 10uA of current through it, the power supply should be 50V! which is not doable for battery-operated chips or wearable devices.  By switching to voltage excitation, 150mV stimulus across that 5M ohm load the resulting current is 30nA. In conclusion, when the impedance is low the current excitation delivers safe and stable current without being skewed by sweat or movement artifacts. When the impedance skyrockets to high values, the mode switches to voltage excitation to avoid compliance rail clipping.  This approach not only resolves the issues faced in BMS3 but also puts the chip safe enough to clinical grade. (Modification on DAC is needed. someone should help designing DAC).

- IEC60601-2-47 timing accuracy compliance demands a low-jitter oscillator with <347 ppm (±11.37 Hz @ 32768 kHz) of frequency accuracy (a designer should help confirm and modify the osc design). MAX30009 by maxim which is clinically graded comes with a PLL to achieve such a high accuracy.